## Slide 1
Quinn Evans · five-week independent ML project · Johns Hopkins PanTS dataset
Research use only — a segmentation and annotation-assist tool, not a diagnostic system.
Notes: 30 min talk + 10 min Q&A. Open by naming the scope out loud: segmentation + a CADe flag, never a diagnosis.

## Slide 2
01
01
The problem and who it's for
Why the pancreas is uniquely hard, and what changes for the person reading the scan.
≈ 4 minutes
Notes: [TRANSIT] 10 seconds. Say the section name and move.

## Slide 3
THE PROBLEM
Why the pancreas is hard
Pancreatic cancer is among the deadliest — outcomes hinge on catching it early
The organ is soft, irregular, and hidden behind other structures
Tumors are often nearly invisible to the human eye on a CT slice
A radiologist reads hundreds of slices per scan, and most scans are normal
10.4%
of scans in this dataset
contain a tumor
A needle-in-a-haystack read, repeated all day.
Problem & business value
1
Notes: [RIFF] Expand on whichever lands: the biology, the invisibility, or the volume of reading. The 10.4% is the business framing — most of the work is ruling things out.

## Slide 4
THE USER
Who it's for, and what changes
Today
Trace the pancreas by hand, slice by slice
Hunt for a lesion that may not be visible
Slow, tedious, and easy to miss something
With this tool
A proposed outline is already drawn
Accept, edit, or reject it
A flag says “look here first”
Export the mask and finish the job
User: radiologist / imaging annotator
Problem & business value
2
Notes: [RIFF] The decision this supports is triage and time: where to look first, and how much tracing work is left.

## Slide 5
SCOPE
What this is — and what it isn't
It IS
A segmentation tool: it outlines the pancreas and any suspicious region
It IS
A CADe assist: it flags “there could be a tumor here” for review
It is NOT
A diagnosis — it never decides whether a patient has cancer
It is NOT
Autonomous — a human reviews and edits every single output
Problem & business value
3
Notes: [TRANSIT] 20 seconds. Stating scope up front pre-empts half the Q&A.

## Slide 6
02
02
Live demo
An unmarked scan, scored live by the deployed model, then checked against the truth.
≈ 7 minutes
Notes: [TRANSIT] Make sure the endpoint is running and the UI is on the Scan library before you advance.

## Slide 7
LIVE DEMO
Three things to watch
1
An unmarked scan
Exactly what a radiologist sees — nothing drawn on it yet
2
The model scores it live
A real call to the deployed endpoint, about half a second
3
The truth is revealed
Expert contours overlaid on the prediction — how close was it?
Every scan shown is held-out data the model never trained on.
Live demo
4
Notes: [TRANSIT] Set expectations in 30 seconds, then switch to the app.

## Slide 8
LIVE DEMO
PanTS Review · localhost:5173 · live inference on port 8000
Live demo
5
Notes: [RIFF] 6-7 minutes. Runsheet: 1. Scan library — real held-out studies 2. Load one UNMARKED — this is what the radiologist sees 3. Analyze — live endpoint call, ~0.6s, 'scored just now' badge 4. Read the finding — flag, size, confidence 5. Reveal source of truth — red vs amber, Dice appears 6. 3D — rotate the pancreas with the lesion inside 7. Show the honest case: the healthy scan it over-calls Fallback: if the endpoint drops, the UI serves cached results — say so out loud, it is a designed behavior.

## Slide 9
LIVE DEMO
What just happened
1
Real model
The registered model, version 1 — the same checkpoint I evaluated
2
Real call
A live HTTP request to a deployed FastAPI endpoint, ~0.6s per scan
3
Real data
Held-out scans from the official test split — never seen in training
Not a recording, not a screenshot.
Live demo
6
Notes: [TRANSIT] 20 seconds. Then move into how it's built.

## Slide 10
03
03
How it's built
The pipeline, the model, and how a checkpoint becomes something a person can use.
≈ 6 minutes
Notes: [TRANSIT] The architecture diagram is next — give yourself room to walk it.

## Slide 11
Notes: [RIFF] 2-3 minutes. Walk it left to right: 1. DATA — 9,901 real volumes on an external drive, indexed, split by patient, audited for broken masks. 2. TRAINING — whole-box preprocessing into one 128 cube, SegResNet fine-tuned from SuPreM, every run in MLflow. 3. EVALUATION + REGISTRY — sliding-window scoring on the untouched test set, then the model is registered as v1. 4. SERVING + INTERFACE — FastAPI endpoint, React/NiiVue front-end, and a human who accepts, edits, or rejects. Then point at the cross-cutting band: that row is the answer to 'where is your DAG'. Ordering, reproducibility, logging and error handling are all delivered — only the scheduler is missing, and the dataset is static. Land on the scope line at the bottom: the model is GIVEN the pancreas region. Say it before anyone asks.

## Slide 12
DATA PIPELINE
The data was the hard part
9,901 real CT volumes — about 380 GB, on an external drive, never in the repo
Split by patient, stratified by tumor, so no patient appears on both sides
Roughly 1 in 8 scans had a broken or empty mask — audited and excluded
Volumes range from 8 to over 1,000 slices, across four scanner manufacturers
9,901
CT volumes indexed
~1 in 8
masks broken or empty
Auditing data integrity turned out to be the work, not a distraction from it.
System walkthrough
7
Notes: [RIFF] Good place to talk about real-world messiness versus clean teaching datasets.

## Slide 13
MODEL
The model
3D SegResNet
Fine-tuned from SuPreM — pretrained by the same Johns Hopkins lab
Three classes: background, pancreas, lesion
Whole-box input: the entire pancreas region as one cube
Trained on Apple Silicon (MPS), ~7-14 h per run
Does the pretraining matter?
I measured it instead of assuming.
0.120
from scratch
0.257
SuPreM transfer
+0.13 lesion Dice — the pretraining is what makes this work at my data scale.
System walkthrough
8
Notes: [RIFF] EXP-09. Single variable, same recipe. From scratch also flagged every healthy scan.

## Slide 14
DEPLOYMENT
From checkpoint to product
MLflow registry
pancreas-lesion-segmenter v1
Checkpoint step 18000
SHA + config + git commit logged
FastAPI endpoint
GET /health, GET /cases
POST /predict → CADe summary
~0.6 s per scan
React + NiiVue UI
Live analyze against the endpoint
Falls back to cached results
3D surfaces from marching cubes
The model is versioned, served, and consumed — not a notebook artifact.
System walkthrough
9
Notes: [TRANSIT] 40 seconds unless asked. This is the layer that makes the demo possible.

## Slide 15
04
04
How I actually worked
Twenty-six experiments, one idea that changed everything, and the day I found my own number was wrong.
≈ 5 minutes
Notes: [TRANSIT] This is the section to modulate. Four stories are coming — lean into whichever the room is leaning into.

## Slide 16
METHOD
26 experiments, bars set in advance
Every experiment: a hypothesis and an accept/reject bar written BEFORE the run
One variable at a time — everything else held fixed
Each result logged with its decision, including the failures
Roughly one experiment per night: a training run is 7 to 14 hours
26
numbered experiments
Accepted, rejected, or null —
all of them written down.
If I couldn't say in advance what would make me abandon an idea, I wasn't testing it.
How I worked
10
Notes: [RIFF] This is the frame for the next three slides. Method first, then the three stories.

## Slide 17
STORY 1
The breakthrough was an idea, not a knob
The model found tumors but over-predicted them — flagging something on 92% of healthy scans
I tried the obvious knobs. Sampling: null. Loss function: null.
Two nulls told me the problem was structural, not a hyperparameter
So: stop feeding random patches. Feed the ENTIRE pancreas region as one cube.
One change
Pancreas Dice
0.72 → 0.81
Specificity
8% → 55%
Lesion Dice
new best
Became the architecture of the final model.
How I worked
11
Notes: [RIFF] The point: the nulls were informative. They told me where the problem wasn't, which is how I found where it was.

## Slide 18
STORY 2
I found out my own number was wrong
An adversarial code audit found my data-scaling script sampled from the wrong column
My training set overlapped my validation set — 30 of the 40 cases I was reporting on had been trained on
I verified it myself, then fixed the root cause rather than the symptom
Added a guard that aborts training if a split ever touches validation — that class of bug can't recur silently
Rebuilt the splits, retrained, and re-reported the honest number
Reported
0.528
Honest
0.474
Lower, real — and it didn't collapse, so the underlying result was genuine.
How I worked
12
Notes: [RIFF] The strongest credibility moment in the talk. Caught before external validation, which is exactly when you want it.

## Slide 19
STORY 3
Two experiments I rejected
Anatomy-aware supervision
Looked like a clear win at an intermediate checkpoint — +0.041 lesion Dice.
I trained both arms to convergence before deciding. The effect vanished. Rejected.
More healthy training data
Cleared its specificity bar decisively — 17% up to 46%.
But detection fell to 88%, three cases below the floor I set in advance. Rejected.
A result you stop at because it looks good is not a result you can trust.
How I worked
13
Notes: [RIFF] Both rejections are wins for the method. The second is the honest 'I didn't move my own goalposts' story.

## Slide 20
05
05
Results, and where it falls short
The honest held-out numbers, what they mean for the user, and the two real limitations.
≈ 5 minutes
Notes: [TRANSIT] Own the limitations here rather than waiting for Q&A.

## Slide 21
RESULTS
Final numbers
Official held-out test set — 901 scans, scored once, never touched during development
96%
detection sensitivity
145 of 151 tumors flagged
0.474
lesion Dice
95% CI 0.42 – 0.52
0.827
pancreas Dice
17%
specificity
on healthy scans
Published PanTS benchmark: ~0.53 lesion Dice (MedFormer 0.529 · R-Super 0.534, which uses additional external data)
Results & limits
14
Notes: [RIFF] Lead with detection — it's the number that matters for a CADe tool. The benchmark line gives the audience a yardstick.

## Slide 22
RESULTS
What those numbers mean for the user
96% detection
Catch rate. Of the tumors present, the tool surfaces almost all of them for review.
0.474 lesion Dice
Edit burden. The outline is a usable starting point that still needs a human pass.
0.827 pancreas Dice
The organ outline is close to expert quality — the tedious tracing is largely done.
17% specificity
False-alarm rate. It cries wolf often, so every flag is a prompt, never a verdict.
Results & limits
15
Notes: [TRANSIT] 45 seconds. Translate once, clearly, then move to the limitations.

## Slide 23
HONEST LIMITS
Limitation: the model is given the pancreas region
It is pointed at the pancreas rather than searching the whole scan on its own.
Why that was the right call for this project
It is a real workflow — in annotation-assist, the reviewer supplies the region
It isolates segmentation quality from localization error, so the number measures one thing
That constraint IS the whole-box breakthrough — the fix for over-prediction came from it
Not ignored: I built and evaluated the autonomous version, with a millimetre-level containment audit
Results & limits
16
Notes: [RIFF] Own it first, then justify. Emphasise: this was a scoped decision, tested, not an oversight.

## Slide 24
HONEST LIMITS
Limitation: specificity is low
It over-calls on healthy scans — 17% specificity is the weakest number in the project.
Deliberately deprioritised — and worked, not forgotten
For a CADe tool, a missed tumor costs more than a false alarm — detection came first
Four documented attempts: sampling ratio, loss function, an anatomical constraint, more healthy data
The model can discriminate — patient-level AUC 0.80 — so it is a threshold problem, not blindness
The honest conclusion: it needs a decision layer in front of it, not a more timid segmenter
Results & limits
17
Notes: [RIFF] The AUC 0.80 line is the sophisticated point — the signal exists, the cut-off is misplaced. Retraining the segmenter to be timid made the outlines worse, which I measured.

## Slide 25
HONEST LIMITS
Where it fails, precisely
Small < 1 cm³
0.067
lesion Dice · 78% detected
Medium 1–8 cm³
0.512
lesion Dice · 100% detected
Large > 8 cm³
0.610
lesion Dice · 100% detected
Small tumors are found but over-drawn — sometimes 25 to 50 times too large. The tool still surfaces them; the outline needs real editing.
Results & limits
18
Notes: [RIFF] Size is the dominant driver. This one behaviour caps Dice AND drives the low specificity — two weaknesses, one cause.

## Slide 26
06
06
Reflection
What I set out to learn in week one, and what five weeks actually taught me.
≈ 3 minutes
Notes: [TRANSIT] Slow down here. This is the part they remember.

## Slide 27
REFLECTION
What I set out to learn
“Learn to set up a 3D image processing pipeline end to end — and prove I can build something that holds up on real-world data.”
Week 1 takeaway
1
The pipeline exists
Indexing through training, evaluation, registry, endpoint, and interface
2
2D → 3D taught me what changes
Memory is the constraint; how you crop is a modelling decision, not preprocessing
3
Real data was the part I underestimated
Broken masks, wild geometry, four scanner makers — auditing WAS the work
Reflection
19
Notes: [RIFF] Close the loop honestly. The 'thrilled if it flags cancer even modestly' line pairs well with 96% detection.

## Slide 28
REFLECTION
How AI actually shaped the work
What I planned
A coding accelerator and an explainer.
What it became
An adversarial reviewer I stopped trusting by default.
The loop: write a spec → a second AI reviews the plan → write code → review the code → a regression test for every bug found
It found my leakage bug — and it had also written that bug
The consistent failure mode: AI is confident about data it has not inspected
So the rule became: make it read the file and show me the output before it asserts anything
Reflection
20
Notes: [RIFF] Honest and non-defensive. Both facts are true: AI found the bug, AI wrote the bug.

## Slide 29
REFLECTION
If I'd had more time
These aren't regrets — they were deferred deliberately. I hit the goals I set for this project.
1
Remove the provided region
Let the system find the pancreas itself, so a raw scan goes in and a result comes out
2
A decision layer for specificity
The AUC says the signal is there — the fix is a gate in front of the model, not a timid model
3
Train on all 9,000 scans
Data scale was the only lever that consistently moved accuracy, and I never found its ceiling
Reflection
21
Notes: [TRANSIT] 30 seconds. Frame as deliberate scope, not regret. One line about carrying it into the capstone is enough.

## Slide 30
Where it landed
96%
of tumors flagged for review
0.474
lesion Dice vs a ~0.53 benchmark
901
held-out scans, scored once
A working 3D pipeline, an honest evaluation, and a tool a radiologist could actually sit down with.
Questions
Notes: [TRANSIT] Land it and stop. 10 minutes of Q&A follows — see the Q&A prep sheet.